## Slide 1: Audible Alert Signal Behavior
QR – Quiet Room
SR – Silent Room
SI – Silent ICU

## Slide 2: Defns. From 60601-1-8
DAS - * DISTRIBUTED ALARM SYSTEM
ALARM SYSTEM that involves more than one item of equipment in a ME SYSTEM intended for delivery
of ALARM CONDITIONS with technical confirmation
NOTE 1 The parts of a DISTRIBUTED ALARM SYSTEM can be widely separated in distance.
NOTE 2 A DISTRIBUTED ALARM SYSTEM is intended to notify OPERATORS of the existence of an ALARM CONDITION.
NOTE 3 For the purposes of this document, technical confirmation means that each element of a DISTRIBUTED ALARM
SYSTEM confirms or guarantees the successful delivery of the ALARM CONDITION to the next element or appropriate
Examples could be:
Traditional Central Station w/o alarm acknowledge;
CDAS - DISTRIBUTED ALARM SYSTEM WITH OPERATOR CONFIRMATION
DISTRIBUTED ALARM SYSTEM that includes the capability to receive an OPERATOR response
Examples could be:
Traditional Central Station; Bed to Bed alarm feature supporting alarm acknowledge; Caregiver worn device supporting alarm acknowledge

## Slide 3: Defns. From 60601-1-8
DIS - * DISTRIBUTED INFORMATION SYSTEM ABOUT ALARM CONDITIONS
system that involves more than one item of equipment in a ME SYSTEM intended to provide
information about ALARM CONDITIONS but does not guarantee delivery of that information
NOTE 1 A DISTRIBUTED INFORMATION SYSTEM ABOUT ALARM CONDITIONS is not intended to notify OPERATORS of the
existence of an ALARM CONDITION as a RISK CONTROL measure. A DISTRIBUTED INFORMATION SYSTEM ABOUT ALARM
CONDITIONS is intended to provide information about an ALARM CONDITION while the OPERATOR is aware of the
existence of the ALARM CONDITION by an ALARM SYSTEM.
NOTE 2 A DISTRIBUTED INFORMATION SYSTEM ABOUT ALARM CONDITIONS is not intended for confirmed delivery of
ALARM CONDITIONS.
Examples could be:
Sometimes referred to as secondary alerting devices: Hallway display of active alarms; Hallway light over room door; Caregiver worn device;

## Slide 4: Defns. From 60601-1-8
EXAMPLE 1 A central station.
EXAMPLE 2 An electronic record-keeping device.
EXAMPLE 3 Remote viewing from home or office.
EXAMPLE 4 Bed-to-bed viewing of ALARM CONDITIONS (e.g. one nurse for two beds).
EXAMPLE 5 Transmission of ALARM CONDITIONS to pagers, cell phones, hand-held computers, etc.
Thus, in a network of bedside PATIENT monitors, one bedside PATIENT monitor can act as a COMMUNICATOR for ALARM CONDITIONS from a different bedside PATIENT monitor. A central station can act as a COMMUNICATOR for ALARM CONDITIONS from multiple PATIENTS. A two-way wireless communication system can act as a COMMUNICATOR for ALARM CONDITIONS to a caregiver in an area far removed from the PATIENT. All these are examples of DISTRIBUTED ALARM SYSTEMS.

## Slide 5: Alert System Variants
Description | Type | Guaranteed Delivery | Operator Confirmation | Example
Aggregates alert signals for devices for a patient’s room | DISr | No | No | Room dashboard, Remote view
 | DASr | Yes | No | 
 | cDASr | Yes | Yes | Room proxy/cockpit
Aggregates alert signals for multiple patients | DISm | No | No | Multi-Pt. hall Dashboard
 | DASm | Yes | No | Multi-Pt. View Station
 | cDASm | Yes | Yes | Multi-Pt. Central Station
Aggregates and directs alert signals to responsible caregivers’ personal devices. | DISp | No | No | Alerts to pager, Mobile viewer
 | DASp | Yes | No | Guaranteed alerts to smartphone
 | cDASp | Yes | Yes | Smartphone with alert ack

## Slide 6: Guide to Table Columns
Location: An example of where this type of use case may be applicable.
Central DISm/DASm/cDASm: Centrally located system that aggregates and signals alerts for Multiple-Patients
  Central DISm – Multi-pt. alert signaling for information only
  Central DASm – Multi-pt. alert signaling (no alert acknowledgement) – primary alert source
  Central cDASm – Multi-pt. alert signaling with alert acknowledgement – primary alert source
Caregiver Location: It may be possible to change system behavior based on the location of the caregiver:
  Unknown
  Out of Room
  In room: in patient room
  At Central Station
  In isolation area (w/o mobile)
  Do not disturb

## Slide 7: Guide to Table Columns
Room DISr/DASr/cDASr: Room located system that aggregates and signals alerts from all networked devices associated with a patient.
  Room DISr – Single-pt. alert signaling for information only
  Room DASr – Single-pt. alert signaling (no alert acknowledgement) – primary alert source
  Room cDASr – Single-pt. alert signaling with alert acknowledgement – primary alert source
Stand-Alone (SA) / Networked (NW):
  Stand-alone: Not connected to a network
  ACM AR: Networked devices communicating alert information using IHE ACM
  SDC NW: Networked devices using IEEE SDC for alert management
Audible:
  Paused: The alert audible signal is automatically temporarily muted in the presence of a ???
  Device Audible: No – Under normal circumstances the individual devices do not generate audible signals
  Room Proxy Audible: No – Under normal circumstances the room proxy does not generate audible signals
  CS Audible: No – Under normal operation the Central Station(s) does not generate audible signals for room alerts
  Mobile Audible: No – Under normal operation a caregiver mobile device generates audible or haptic signals for room alerts
Quiet Room (QR) / Silent Room (SR) / Silent ICU (SI):

## Slide 8: No Multi-Pt. Central Alert Aggregator
ICE / Clinical Workspace
SDC / ACM
DISr / DASr / cDASr
Device
Device
Device
SDC / ACM
DISr / DASr / cDASr
Device
Device
Some devices NW’d
All devices NW’d
ICE / Clinical Workspace
SDC / ACM
Device
Device
Device
Device
No Alert Aggregator
Room
Device

## Slide 9: No Multi-Pt. Central Alert Aggregator
Location | Central | Caregiver Location | Room | Stand-alone / Networked | Audible |  |  |  | QR /SR /SI
 |  |  |  |  | @Device | @Room Proxy | @CS | @Mobile | 
OR | None | In room | No / DISr | SA & NW | Yes | N/A | N/A | N/A | --
 |  |  |  | ACM AR | Yes | N/A | N/A | N/A | --
 |  |  |  | SDC NW | Yes | N/A | N/A | N/A | --
 |  |  | DASr / cDASr | SA & NW | SA -Yes NW - No | NW - Yes | N/A | N/A | --
 |  |  |  | ACM AR | No | Yes | N/A | N/A | --
 |  |  |  | SDC NW | No | Yes | N/A | N/A | --

## Slide 10: Central Multi-Patient Alert Aggregator DISm and/or DISp
DISm
SDC / ACM
SDC / ACM
DISr / DASr / cDASr
Device
Device
Some devices NW’d
All devices NW’d
Room
SDC / ACM
Device
Device
Device
Device
No Alert Aggregator
Room
Device
SDC / ACM
DISr / DASr / cDASr
Device
Device
Room
Device
DISp

## Slide 11: Central Multi-Patient Alert Aggregator DISm
Location | Central DAS / DIS | Caregiver Location | Room DAS Proxy | Stand-alone / Networked | Audible |  |  |  | QR /SR /SI
 |  |  |  |  | @Device | @Room Proxy | @CS | @Mobile | 
MedSurg | DISm | Unknown / Out of Rm | No / DISr | SA & NW | Yes | N/A | NW - Yes | N/A | --
 |  |  |  | ACM AR | Paused | N/A | Yes | N/A | QR
 |  |  |  | SDC NW | Paused | N/A | Yes | N/A | QR
 |  |  | DASr / cDASr | SA & NW | SA -Yes NW – No | Paused | Yes | N/A | --
 |  |  |  | ACM AR | Paused | Paused | Yes | N/A | QR
 |  |  |  | SDC NW | No | Paused | Yes | N/A | QR
 |  | In room | No / DISr | SA & NW | Yes | N/A | No | N/A | --
 |  |  |  | ACM AR | Yes | N/A | No | N/A | --
 |  |  |  | SDC NW | Yes | N/A | No | N/A | --
 |  |  | DASr / cDASr | SA & NW | SA -Yes NW – No | NW – Yes | No | N/A | --
 |  |  |  | ACM AR | Paused | Yes | No | N/A | --
 |  |  |  | SDC NW | No | Yes | No | N/A | --

## Slide 12: Central Multi-Patient Alert Aggregator DISp
Location | Central DAS / DIS | Caregiver Location | Room DAS Proxy | Stand-alone / Networked | Audible |  |  |  | QR /SR /SI
 |  |  |  |  | @Device | @Room Proxy | @CS | @Mobile | 
MedSurg | DISp | Unknown / Out of Rm | No / DISr | SA & NW | Yes | N/A | N/A | NW - Yes | --
 |  |  |  | ACM AR | Paused | N/A | N/A | Yes | QR
 |  |  |  | SDC NW | Paused | N/A | N/A | Yes | QR
 |  |  | DASr / cDASr | SA & NW | SA -Yes NW – No | Paused | N/A | Yes | --
 |  |  |  | ACM AR | Paused | Paused | N/A | Yes | QR
 |  |  |  | SDC NW | No | Paused | N/A | Yes | QR
 |  | In room | No / DISr | SA & NW | Yes | N/A | N/A | No | --
 |  |  |  | ACM AR | Yes | N/A | N/A | No | 
 |  |  |  | SDC NW | Yes | N/A | N/A | No | QR
 |  |  | DASr / cDASr | SA & NW | SA -Yes NW – No | NW – Yes | N/A | No | --
 |  |  |  | ACM AR | Paused | Yes | N/A | No | 
 |  |  |  | SDC NW | No | Yes | N/A | No | QR

## Slide 13: Central Multi-Patient Alert Aggregator DASm and/or DASp
DASm
SDC / ACM
SDC / ACM
DISr / DASr / cDASr
Device
Device
Some devices NW’d
All devices NW’d
Room
SDC / ACM
Device
Device
Device
Device
No Alert Aggregator
Room
Device
SDC / ACM
DISr / DASr / cDASr
Device
Device
Room
Device
DASp

## Slide 14: Central Multi-Patient Alert Aggregator DASm
Location | Central DAS / DIS | Caregiver Location | Room DAS Proxy | Stand-alone / Networked | Audible |  |  |  | QR /SR /SI
 |  |  |  |  | @Device | @Room Proxy | @CS | @Mobile | 
ICU | DASm | Unknown / Out of Rm | No / DISr | SA & NW | SA -Yes NW – No | N/A | SA – No NW - Yes | N/A | --
 |  |  |  | ACM AR | Paused | N/A | Yes | N/A | 
 |  |  |  | SDC NW | No | N/A | Yes | N/A | SR
 |  |  | DASr / cDASr | SA & NW | SA -Yes NW – No | No | SA - No NW - Yes | N/A | --
 |  |  |  | ACM AR | Paused | Paused | Yes | N/A | 
 |  |  |  | SDC NW | No - Ack | Paused? | Yes | N/A | SR
 |  | In room | No / DISr | SA & NW | Yes | N/A | SA – No NW - Paused | N/A | --
 |  |  |  | ACM AR | Yes | N/A | Paused | N/A | 
 |  |  |  | SDC NW | Yes | N/A | No | N/A | SR
 |  |  | DASr / cDASr | SA & NW | SA -Yes NW – No | Yes | SA - No NW - Paused | N/A | --
 |  |  |  | ACM AR | Paused | Yes | Paused | N/A | 
 |  |  |  | SDC NW | No - Ack | Yes | No | N/A | SR

## Slide 15: Central Multi-Patient Alert Aggregator DASp
Location | Central DAS / DIS | Caregiver Location | Room DAS Proxy | Stand-alone / Networked | Audible |  |  |  | QR /SR /SI
 |  |  |  |  | @Device | @Room Proxy | @CS | @Mobile | 
ICU | DASp | Unknown / Out of Rm | No / DISr | SA & NW | SA -Yes NW – No | N/A | N/A | SA – No NW - Yes | --
 |  |  |  | ACM AR | Paused | N/A | N/A | Yes | 
 |  |  |  | SDC NW | No | N/A | N/A | Yes | SR
 |  |  | DASr / cDASr | SA & NW | SA -Yes NW – No | No | N/A | SA - No NW - Yes | --
 |  |  |  | ACM AR | Paused | Paused | N/A | Yes | 
 |  |  |  | SDC NW | No - Ack | Paused? | N/A | Yes | SR
 |  | In room | No / DISr | SA & NW | Yes | N/A | N/A | SA – No NW - Paused | --
 |  |  |  | ACM AR | Yes | N/A | N/A | Paused | 
 |  |  |  | SDC NW | Yes | N/A | N/A | No | SR
 |  |  | DASr / cDASr | SA & NW | SA -Yes NW – No | Yes | N/A | SA - No NW - Paused | --
 |  |  |  | ACM AR | Paused | Yes | N/A | Paused | 
 |  |  |  | SDC NW | No - Ack | Yes | N/A | No | SR

## Slide 16: Central Multi-Patient Alert Aggregator DASm & DASp
Location | Central DAS / DIS | Caregiver Location | Room DAS Proxy | Stand-alone / Networked | Audible |  |  |  | QR /SR /SI
 |  |  |  |  | @Device | @Room Proxy | @CS | @Mobile | 
ICU | DASm & DASp | Unknown / Out of Rm | No / DISr | SA & NW | SA -Yes NW – No | N/A | SA – No NW - Yes | SA – No NW - Yes | --
 |  |  |  | ACM AR | Paused | N/A | Yes | Yes | 
 |  |  |  | SDC NW | No | N/A | Yes | Yes | SR
 |  |  | DASr / cDASr | SA & NW | SA -Yes NW – No | No | SA - No NW - Yes | SA - No NW - Yes | --
 |  |  |  | ACM AR | Paused | Paused | Yes | Yes | 
 |  |  |  | SDC NW | Paused? | Paused? | Yes | Yes | SR
 |  | In room | No / DISr | SA & NW | Yes | N/A | SA – No NW - Paused | SA – No NW - Paused | --
 |  |  |  | ACM AR | Yes | N/A | Paused | Paused | 
 |  |  |  | SDC NW | Yes | N/A | No | No | SR
 |  |  | DASr / cDASr | SA & NW | SA -Yes NW – No | Yes | SA - No NW - Paused | SA - No NW - Paused | --
 |  |  |  | ACM AR | Paused | Yes | Paused | Paused | 
 |  |  |  | SDC NW | No - Ack | Paused? | No | No | SR

## Slide 17: Central Multi-Patient Alert Aggregator cDASm and/or cDASp
cDASm
SDC / ACM
SDC / ACM
DISr / DASr / cDASr
Device
Device
Some devices NW’d
All devices NW’d
Room
SDC / ACM
Device
Device
Device
Device
No Alert Aggregator
Room
Device
SDC / ACM
DISr / DASr / cDASr
Device
Device
Room
Device
cDASp

## Slide 18: Central Multi-Patient Alert Aggregator cDASm
Location | Central DAS / DIS | Caregiver Location | Room DAS Proxy | Stand-alone / Networked | Audible |  |  |  | QR /SR /SI
 |  |  |  |  | @Device | @Room Proxy | @CS | @Mobile | 
ICU | cDASm | Unknown / Out of Rm | No / DISr | SA & NW | SA -Yes NW – No | N/A | SA – No NW - Yes | N/A | --
 |  |  |  | ACM AR | Paused | N/A | Yes | N/A | 
 |  |  |  | SDC NW | No | N/A | Yes | N/A | SR
 |  |  | DASr / cDASr | SA & NW | SA -Yes NW – No | N/A | SA - No NW - Yes | N/A | --
 |  |  |  | ACM AR | Paused | Paused | Yes | N/A | 
 |  |  |  | SDC NW | No - Ack | No - Ack | Yes | N/A | SR
 |  | In room | No / DISr | SA & NW | Yes | N/A | SA – No NW - Paused | N/A | --
 |  |  |  | ACM AR | Yes | N/A | Paused | N/A | 
 |  |  |  | SDC NW | Yes | N/A | No | N/A | SR
 |  |  | DASr / cDASr | SA & NW | SA -Yes NW – No | Yes | SA - No NW - Paused | N/A | --
 |  |  |  | ACM AR | Paused | Yes | Paused | N/A | 
 |  |  |  | SDC NW | No - Ack | Yes | No | N/A | SR

## Slide 19: Central Multi-Patient Alert Aggregator cDASp
Location | Central DAS / DIS | Caregiver Location | Room DAS Proxy | Stand-alone / Networked | Audible |  |  |  | QR /SR /SI
 |  |  |  |  | @Device | @Room Proxy | @CS | @Mobile | 
ICU | cDASp | Unknown / Out of Rm | No / DISr | SA & NW | SA -Yes NW – No | N/A | N/A | SA – No NW – Yes | --
 |  |  |  | ACM AR | Paused | N/A | N/A | Yes | 
 |  |  |  | SDC NW | No | N/A | N/A | Yes | SR
 |  |  | DASr / cDASr | SA & NW | SA -Yes NW – No | N/A | N/A | SA - No NW - Yes | --
 |  |  |  | ACM AR | Paused | Paused | N/A | Yes | 
 |  |  |  | SDC NW | No - Ack | No - Ack | N/A | Yes | SR
 |  | In room | No / DISr | SA & NW | Yes | N/A | N/A | SA – No NW - Paused | --
 |  |  |  | ACM AR | Yes | N/A | N/A | Paused | 
 |  |  |  | SDC NW | Yes | N/A | N/A | No | SR
 |  |  | DASr / cDASr | SA & NW | SA -Yes NW – No | Yes | N/A | SA - No NW - Paused | --
 |  |  |  | ACM AR | Paused | Yes | N/A | Paused | 
 |  |  |  | SDC NW | No - Ack | Yes | N/A | No | SR

## Slide 20: Central Multi-Patient Alert Aggregator cDASm & cDASp
Location | Central DAS / DIS | Caregiver Location | Room DAS Proxy | Stand-alone / Networked | Audible |  |  |  | QR /SR /SI
 |  |  |  |  | @Device | @Room Proxy | @CS | @Mobile | 
ICU | cDASm & cDASp | Unknown / Out of Rm | No / DISr | SA & NW | SA -Yes NW – No | N/A | SA – No NW – Yes | SA – No NW – Yes | --
 |  |  |  | ACM AR | Paused | N/A | Yes | Yes | 
 |  |  |  | SDC NW | No | N/A | Yes | Yes | SR
 |  |  | DASr / cDASr | SA & NW | SA -Yes NW – No | N/A | SA - No NW - Yes | SA - No NW - Yes | --
 |  |  |  | ACM AR | Paused | Paused | Yes | Yes | 
 |  |  |  | SDC NW | No - Ack | No - Ack | Yes | Yes | SR
 |  | In room | No / DISr | SA & NW | Yes | N/A | SA – No NW - Paused | SA – No NW - Paused | --
 |  |  |  | ACM AR | Yes | N/A | Paused | Paused | 
 |  |  |  | SDC NW | Yes | N/A | No | No | SR
 |  |  | DASr / cDASr | SA & NW | SA -Yes NW – No | Yes | SA - No NW - Paused | SA - No NW - Paused | --
 |  |  |  | ACM AR | Paused | Yes | Paused | Paused | 
 |  |  |  | SDC NW | No - Ack | Yes | No | No | SR

## Slide 22
Typical Alert Distribution System network topology today

## Slide 23: Current 2020 - Typical Network Architecture
Monitor Central Station
Alert Distribution
System
Room
Proprietary
Bedside Monitor
Device
Device
Device
Device Central Station
Room
Bedside Monitor
Device
Device
Device
Room
Bedside Monitor
Device
Device
Device
Monitor GW
Device GW
Proprietary
Typical Room has:
Networked bedside
  Bedside networked to Central Cockpit
  Device such as vent connected to bedside
Networked infusion pumps connected to pump central
Other devices – standalone (not networked)
Gateways that convert proprietary protocols to PCD ACM (DEC, etc.)
ACM

## Slide 24: Current 2020 - Typical Communication Paths
Monitor Central Station
Alert Distribution
System
Room
Bedside Monitor
Device
Device
Device
Device Central Station
Room
Bedside Monitor
Device
Device
Device
Room
Bedside Monitor
Device
Device
Device
Monitor GW
Device GW
ACM

## Slide 25
Step a:
SDPi implementation for Quiet Hospital
Mix of stand-alone and SDC devices – Partial QH

## Slide 26: Target 202a? - SDPi Quiet Hospital Network Architecture
Central Cockpit
Alert Distribution
System
ICE / Clinical Workspace
SDC
SDC
Cockpit
Smart Alarms Apps
Proxy
Device
Device
Device
SDC
Cockpit Smart Alarms Apps
Device
Device
Device
SDC
Device
Device
Device
Device
SOMDS w/proxy
SOMDS no proxy
SOMDS no cockpit
SDC to PCD ACM Gateway
Assumes devices in room are connected via SDC or are standalone
Quiet Hospital concept achieved with connected devices.
Nurse devices can receive and acknowledge alerts via the Alert Distribution System
ICE / Clinical Workspace
Room
SDC
ACM

## Slide 27: Target 202a? - SDPi Quiet Hospital Network Architecture
Central Cockpit
Alert Distribution
System
SDC
Cockpit
Smart Alarms Apps
Proxy
Device
Device
Device
SDC
Cockpit
Smart Alarms Apps
Device
Device
Device
SDC
Device
Device
Device
Device
ACM
SDC to PCD ACM Gateway
ICE / Clinical Workspace
SOMDS w/proxy
SOMDS no proxy
SOMDS no cockpit
ICE / Clinical Workspace
Room

## Slide 28
Step b:
SDPi implementation for Quiet Hospital
Only SDC devices – full QH

## Slide 29: Target 202b? - SDPi Quiet Hospital Network Architecture
Central Cockpit
Alert Distribution
System
SDC
SDC
SDC
Cockpit
Smart Alarms Apps
Proxy
Device
Device
Device
SDC
Cockpit
Smart Alarms Apps
Device
Device
Device
SDC
Device
Device
Device
Device
ACM
SDC to PCD ACM Gateway
Assumes devices in room are all connected via SDC
Quiet Hospital concept achieved
Nurse devices can receive and acknowledge alerts via the Alert Distribution System
ICE / Clinical Workspace
SOMDS w/proxy
SOMDS no proxy
SOMDS no cockpit
ICE / Clinical Workspace
Room

## Slide 30: Target 202b? - SDPi Quiet Hospital Network Architecture
Central Cockpit
Alert Distribution
System
SDC
Cockpit
Smart Alarms Apps
Proxy
Device
Device
Device
SDC
Cockpit
Smart Alarms Apps
Device
Device
Device
SDC
Device
Device
Device
Device
ACM
SDC to PCD ACM Gateway
ICE / Clinical Workspace
SOMDS w/proxy
SOMDS no proxy
SOMDS no cockpit
ICE / Clinical Workspace
Room

## Slide 31
Step c:
SDPi implementation for Silent Hospital
Nurse devices can control bedside SDC devices

## Slide 32: Target 202c? - SDPi Silent Hospital Network Architecture
Central Cockpit
Alert Distribution
System
SDC
SDC
SDC
Cockpit
Smart Alarms Apps
Proxy
Sensor Proxy
Device
Device
SDC
Cockpit
Smart Alarms Apps
Sensor Proxy
Device
Device
SDC
Sensor Proxy
Device
Device
Device
ACM
Assumes Alert Distribution System can talk SDC natively, thereby eliminating the need to a Gateway.
Gateway may still be required for Device to EHR applications.
Nurse alert devices connect directly to SDC network.
Nurse alert devices can fully manage device alerts and exposed settings.
Can now fully support Silent Hospital concept.
ICE / Clinical Workspace
SOMDS w/proxy
SOMDS no proxy
SOMDS no cockpit
ICE / Clinical Workspace
Room

## Slide 33: Target 202c? - SDPi Silent Hospital Communication Paths
Central Cockpit
Alert Distribution
System
SDC
Cockpit
Smart Alarms Apps
Proxy
Sensor
Proxy
Device
Device
Cockpit
Smart Alarms Apps
SDC
Sensory Proxy
Device
Device
SDC
Sensor Proxy
Device
Device
Device
Not all communication paths drawn
ICE / Clinical Workspace
SOMDS w/proxy
SOMDS no proxy
SOMDS no cockpit
ICE / Clinical Workspace
Room

## Slide 34
Step d:
SDPi implementation for Silent Hospital
SDC replaces ACM

## Slide 35: Target 202d? - SDPi Silent Hospital Network Architecture
SDC
SDC
SDC
Cockpit
Smart Alarms Apps
Proxy
Sensor Proxy
Device
Device
SDC
Cockpit
Smart Alarms Apps
Sensor Proxy
Device
Device
SDC
Sensor Proxy
Device
Device
Device
ACM or FHIR or SDC
Assumes Alert Distribution System can talk SDC natively, thereby eliminating the need to a Gateway.
Gateway may still be required for Device to EHR applications.
Nurse alert devices connect via:
SDC or ACM (v2) or FHIR or some combination thereof…
ICE / Clinical Workspace
SOMDS w/proxy
SOMDS no proxy
SOMDS no cockpit
ICE / Clinical Workspace
Room
Central Cockpit
Alert Distribution
System

## Slide 36: Target 202e? - SDPi Silent Hospital Network Architecture
Central Cockpit
Alert Distribution
System
SDC
SDC
Cockpit
Smart Alarms Apps
Proxy
Sensor Proxy
Device
Device
SDC
Cockpit
Smart Alarms Apps
Sensor Proxy
Device
Device
SDC
Sensor Proxy
Device
Device
Device
All alert communication and related management is via SDC.
Where does FHIR fit in?
ICE / Clinical Workspace
SOMDS w/proxy
SOMDS no proxy
SOMDS no cockpit
ICE / Clinical Workspace
Room
SDC
FHIR Server

## Slide 37: Target 202d? - SDPi Silent Hospital Communication Paths
Central Cockpit
Alert Distribution
System
SDC
SDC
SDC
Cockpit
Smart Alarms Apps
Proxy
Sensor Proxy
Device
Device
Cockpit
Smart Alarms Apps
SDC
Sensor Proxy
Device
Device
SDC
Sensor Proxy
Device
Device
Device
Not all communication paths drawn
ICE / Clinical Workspace
SOMDS w/proxy
SOMDS no proxy
SOMDS no cockpit
ICE / Clinical Workspace
Room