## Slide 1: Safe, Effective & Secure Device SDC-SDPi+FHIR Interoperability
FHIR is a trademark of Health Level 7, International.
SDC is a registered trademark of OR.NET
Establishing a framework for
Trusted Interoperable Product Decoupling
Using ISO/IEEE 11073 SDC / IHE SDPi / HL7 FHIR + ISO/IEC JWG7 SES Standards

## Slide 2: Safe, Effective & Secure MDI using SDC-SDPi+FHIR
The Question
Journey from SDC to SDPi+FHIR …
Overview of SDC & SDPi for MDI
SES MDI – Proposed “Getting Started” Roadmap
From Narratives to Plug-n-Trust

## Slide 3: SES SDC-SDPi Device Interoperability
Problem: Medical device interoperability (MDI) standards & Medical Technology Safety, Effectiveness & Security (SES) standards exist in parallel universes BUT products allowed for patient use must meet both the informatics interoperability technology requirements + quality, regulatory, and legal requirements.
Question: Can a framework be created to enable
          Trusted Interoperable Product Decoupling
      Using
      ISO/IEEE 11073 SDC, IHE SDPi & HL7 FHIR Interoperability Standards
      +
      ISO/IEC JWG7 Safety, Effectiveness & Security Standards?
NOTE: MDI standards align with ISO/TC 215 WG2; SES with ISO/IEC JWG7

## Slide 4: Acronyms
SES Safety, Effectiveness & Security – A handy way to refer to the JWG7 focus and family of standards
SDC Service-oriented Device Connectivity – Family of ISO/IEEE/CEN 11073 device interoperability standards (see TC215 WG2)
DPI Device Point-of-care Interoperability – IHE Devices sub-group designation
SDPi Service-oriented Device Point-of-care Interoperability – IHE Devices profiles using SDC & FHIR (and HL7 V2 …)
MDI Medical Device Interoperability – General label covering entire subject area
FHIR Fast Healthcare Interoperability Resources – General purpose resource-based health information exchange standards typically using RESTful architecture
QRL Quality / Regulatory / Legal stakeholder communities – For lack of a better TLA!
IHE Integrating the Healthcare Enterprise – An international standards “profiling” organization (ISO/TC 215 Liaison A)
JWG7 Joint Working Group between ISO/TC 215 & IEC/SC 62A – Home of the SES family of standards

## Slide 5: A word about “Safety, Effectiveness & Security”
Two+ ISO/IEC standards & guidance documents that address SES MDI across the life cycle of health software, Health I.T. & Incorporated devices.
Primary focus area of SES @ SDPi+FHIR white paper project
SES
MDI Trust Gap!
Diagram from 81001-1:2020 under Open Government License: see https://www.nationalarchives.gov.uk/doc/open-government-licence/version/3/
As explained to the SDC-SDPi+FHIR MDI community …

## Slide 6
MDI Journey so far …
… recent history

## Slide 7: ISO/IEEE 11073 SDC – 15 Year Journey
2004
BMBF Vision SOMIT FUSION / OrthoMIT
Foundation for the idea of interoperability
2004
2010
TekoMed
Feasibility study to prove the SOA approach for medical devices
Dienst-Orientierte OP Integration (DOOP)
Networking project with various medical vendors to implement DPWS and demonstrate interoperability
2011
2013
BMBF-OR.NET
A project funded by the German Ministry of Education and Research to consolidate all medical device interoperability research activities in Germany
2015
OR.NET Consortium
An association of different stakeholders in medical device interoperability
IEEE 11073-20701 Standard approved
Service Oriented Medical Device Exchange Architecture & Protocol Binding
2018
2016
IEEE 11073-20702 Standard approved
Medical Devices Communication Profile for Web Services
IEEE 11073-10207 Standard approved
Domain Information and Service Model for Service-Oriented Point-of-Care Medical Device Communication
2017
NOTE: This roughly parallels the timelines for IHE Devices Domain & HL7 Devices WG

## Slide 8: IHE SDPi – Building on the ISO/IEEE & HL7 Core
https://www.ihe.net/uploadedFiles/Documents/PCD/IHE_PCD_WP_SDPi_Rev1-1_Pub_2019-11-01.pdf
https://wiki.ihe.net/index.php/SDC@IHE_White_Paper
SDC@IHE 2019 Initiative laid the foundation for IHE
  Service-oriented Device Point-of-Care Interoperability (SDPi)
profile family for “PRACtical” device-to-device interoperability

## Slide 9: 2020 Joint HL7-IHE Gemini Project
Full slide deck @ confluence.hl7.org “Device Interoperability using SDPi+FHIR” page

## Slide 10: Gemini Project Deliverables & Governance
One set of cohesive, coordinated deliverables
  IHE SDPi Supplement – published 2020 JUL, PAT/CAT testing Q3/4 ‘20
  HL7 DoF IG supporting SDC integration & Alerting – ballot in 2020
    DoF IG (proposed) for Device Information Consumers (title TBD)
  Joint White Papers:
    “What is a device?” - including AI/ML SAMD, across use context geographies
    “Safe, Effective & Secure MDI Using SDC/SDPi + FHIR” – Quality / Regulatory / Legal Considerations
    "Accelerating Safe, Effective and Secure Remote Monitoring and Mobile Health Interoperable Solutions“ – How do you know that a rapid response to address crisis (e.g., pandemic) challenges is safe enough, effective enough, and secure enough to allow for implementation & use?
Governance based on HL7 or IHE project home organization processes

## Slide 11
A bit more on …
SDC-SDPi+FHIR
… for MDI

## Slide 12
Standards for device-to-device interoperability at Point-of-Care ?
Gateway
Alarm Delegation Services
EMR
LAB
PACS
Central Station
Enterprise
Point-of-Care
Patient Monitoring
Bedside information hub
Ventilation
Infusion systems
imaging systems

## Slide 13: What is SDC?
Service-oriented Device Connectivity
A family of international standards for interoperable exchange of real-time information between medical devices and external systems in dynamic IP networks

## Slide 14: SDC Basics …
SDC Point-of-Care Context
HIT “Enterprise” Context
Device-to-Device
Plug-and-Play for
Reporting / Alerting & Controlling
(PRACtical Interoperability)

## Slide 15: IEEE 11073 SDC Standards “Cathedral”
Device Specializations
Participant Key
Purposes
Core Connectivity
Standards
Nomenclature

## Slide 16: Example: Quiet Hospital / Silent Hospital Narratives
NOTE: Current version of these narrative documents are available on the Gemini SDPi+FHIR Quiet Hospital Confluence pages.

## Slide 17: SDC & Quiet Hospital/Silent ICU @ HIMSS’20 …
~ 80 - 95 % clinically irrelevant
50 % are not noticed
up to 40 min. to alarm confirmation
~ 40 different sounds in one ICU
Tomorrow
Z
Z
Z
ok!
road to
1
alarming device per patient
Reduction of alarms
Alarm Distribution
Today
SILENT ICU BY ALARM SIGNAL DELEGATION

## Slide 18: SDC & Quiet Hospital/Silent ICU @ HIMSS’20 …
The alarm producer has to make all information available that are necessary for the remote alarm notifiers, like alert condition presence, alert manifestation, etc. Interoperability and semantical interpretability have to be ensured.
The system has to be suitable for multiple alarm producers and several remote alarm notifying devices.
The alarm producer has to be able to determine whether other devices are ready to generate the alarm notification.
The alarm producer has to be able to observe that the alert is generated correctly.
Some more information: „A Safe and Interoperable Distributed Alarm Notification System for PoC Medical Devices using IEEE 11073 SDC”, Kasparick et al.
“Delegation” – Safely enabling one system to annunciate alerts on the behalf of another system
SILENT ICU BY ALARM SIGNAL DELEGATION - Requirements

## Slide 19: Alert Signal Delegation Use Case
From IEEE 11073-10207 (BICEPS)
ALERT SIGNAL DELEGATION is the capability of a POC MEDICAL DEVICE to let another PARTICIPANT generate a POC MEDICAL DEVICE’s ALERT SIGNAL as primary ALERT SIGNAL in order to remotely indicate the presence of an ALERT CONDITION on the POC MEDICAL DEVICE.
a POC MEDICAL DEVICE delegates its ALERT SIGNAL generation to another PARTICPANT, e.g., to facilitate a silent workplace
Delegable & Fallback Alert Signals

## Slide 20: Alert Signal Distribution Combining IEEE 11073 SDC & IHE PCD ACM
Patient Monitoring
Ventilation
IEEE 11073 SDC
IHE PCD ACM
PCD-04/05
PCD-06/07
Gateway
Alert
Manager
Alert Communicator
Alert
Provider
Alert
Provider
Alert
Consumer
Alert
Reporter
SDPi Alerting
(SDPi-A)

## Slide 21: SDC – Security is a Core Capability
SDC Communication
Standard based encryption of communication: TLS 1.2+
Each endpoint has a certificate.
Authentication at both endpoints
No communication if no trust chain can be established
Certificates are used to secure communication
Authorization and Authentication
Certificates carry roles of participants
Each device can decide if remote control is OK based on certificate roles and certifying organization
Ventilator V1
Vendor: xyz Corp
Roles:
Metric Provider
Operation Provider
Certified by:
Vitals Dashboard D1
Vendor: ABC Corp
Roles:
Metric Consumer
Operation Consumer
Certified by:
SDC = Enables Trusted Interoperable Product Decoupling
(Stretch goal: Computable IFU discoverable at Plug-n-Trust network connection & configuration … just sayin’ …)

## Slide 22: SDPi+FHIR Technical Concept
The Amsterdam Sessions 2019.11
  IHE Devices domain Technical Framework Approach Established
  IHE SDPi Supplement w/ 4 profiles under development – in parallel
  Will include content for TF-1 -2 & -3
  IHE Connectathon testing planned for 2020

## Slide 23: IHE SDPi – Building on the ISO/IEEE & HL7 Core
IHE SDPi Profiles for “PRACtical” Device Interoperability
SDPi Plug-and-Trust Connectivity
(SDPi-P)
SDPi external Control
(SDPi-xC)
SDPi Reporting
(SDPi-R)
SDPi Alerting
(SDPi-A)
“medical” interoperability purpose
HL7 FHIR
PoCD IG
FHIR Base Resources
SDC PKP Standards
SDC Core Standards
DEC
HL7 V2 Messaging
ACM
IHE “Gateway” Actors Defined

## Slide 24: SDC PKP’s – Pathway to Trust at the Interface
SDC SES Design Approach
  Establish Key Purposes of Interoperability from MDS to interface:
    Connectivity
    Reporting
    Alerting
    External Control
  Integrate “testable assertions” into each standard (”Rxxxx”)
  Bind Rxxxx from layer to layer, incl. risks from one layer mapped into RCMs in the next
  Address both Functional & Non-functional (RM) Aspects
  Per-device System Function Contribution defined
BICEPS / Dev Specializations
MDPWS/WS-*

## Slide 25: SDC Conformance Principles
OR.NET* white paper lays foundation for traceability from PKPs to Conformity Assessment (CA) and certified safe-effective-secure (SES) interoperable medical device system components.
IEEE 11073 -10700 Base PKP standard includes annex with summary of Conformance Principles
* Germany-based non-profit academic / industry group

## Slide 26: SES Plug-n-Trust Requires Rigorous Testing
Leverage the established Device Interoperability tooling …
Objective: Leverage IHE Test & Tooling to establish an SDC-enabled interoperable medical technology ecosystem where certified test reports can be directly included in regulatory submissions
NIST Device Tools

## Slide 27: For more general SDC-SDPi+FHIR Info …
Primary source: confluence.hl7.org “SDPi+FHIR Community”
Schlichting / Unity presentation @ https://bit.ly/Record-MDI
HIMSS’20 Klotz/Cooper SDC MDI digital session (to be posted)
IHE DEV SDPi White Paper –
    Wealth of info including Bibliography with useful
    links to papers and presentations and open source
    Implementations, etc.
https://www.ihe.net/uploadedFiles/Documents/PCD/IHE_PCD_WP_SDPi_Rev1-1_Pub_2019-11-01.pdf

## Slide 28
Connecting the dots from
… inspiring narratives
… to Plug-n-Trust!

## Slide 29
Full slide deck @ confluence.hl7.org “SDPi+FHIR Tools“ page

## Slide 30: From Inspiring Narratives to Plug-n-Trust
Establishing a framework strategy for …
    Capturing detailed requirements from use narratives in a clear and computable manner;
    Enabling mapping & traceability & simulation through each technology and quality specification layer;
    Testing V&V conformity assessment at the interoperable component and system of interoperable products level
… an ecosystem of trusted, safe, effective & secure, plug-and-play device & health software solutions using
SDC-SDPi + FHIR

## Slide 31: Journey starts @ Inspiring Narratives
https://wiki.ihe.net/index.php/SDC@IHE_White_Paper
Gemini SES MDI Project:
  Functional Endoscopic Sinus Surgery / Operating Room Integration
  Quiet Hospital / Silent ICU
  JHU/APL MDIRA/ICE for Military & Disaster Casualty Care
  Preeclampsia During Pregnancy – Home to Clinic to Hospital
  (in process) “Loopers” Type 1 Diabetes – Person-crafted Closed-Loop Control
Gemini SES MDI Use Case home pages @
HL7 Confluence “Device Interoperability using SDPi+FHIR”/SDPi+FHIR Use Cases

## Slide 32: IHE Use Case Analysis – Hierarchy of Concepts
Source: Journal of AHIMA, 2017 July – October series, “Understanding User Needs for Interoperability: Standards for Use Cases in eHealth”

## Slide 33: “SDPi+FHIR” – Introduced @ HIMSS’20 …

## Slide 34: Expanding the Requirements-to-Interfaces Model
Note: model is not perfect … but very useful!
Multiple Standards / Specifications
Multiple Organizations / SDOs
Multiple Projects / Initiatives
ISO/IEEE 11073 SDC –
“Cathedral” Model

## Slide 35: Charting the course from narratives to interfaces
Traceability from Narrative to
Plug-n-Trust Interface!
Each “layer” specifies requirements to be mapped to the next
Each “layer” adds its own set of requirements
Requirements align with safety, effectiveness & security (SES)
NOTE: Null layers allowed but generally required to achieve the SES Plug-n-Trust objective.
Comprehensive System Function Contribution (SFC) Plug-n-Trust Interfaces
/ Ecosystem

## Slide 36: Charting the course from narratives to interfaces
The Challenge …
  Establish computable traceability from top to bottom (plug-n-trust) interface
  Automate mapping and testability
  Scale from simple to complex
  Layer-specific and layer-to-layer (mapped) requirements representation
  Lower the bar for specification creation, adoption, implementation and use
  Maximize use of off-the-shelf open, “free”, widely adopted and used tooling
  Ensure integration with corporate enterprise-grade tooling (including DOORS, Enterprise Architect, MagicDraw, etc.)
  Support V-model CA and 3rd party (17025) “test houses”
  …
“Mapping” remains the hardest nut to crack in this list of challenges. Especially across multiple standards, organizations and projects!

## Slide 37: Capturing Computable Use Narrative Requirements
Question: How to capture computable, testable requirements starting with user narratives?
User Narratives / Use Cases / Requirements
Narrative “User Story” Text
Gherkin / Cucumber Representation
ReqIF (Requirements Interchange Format) Representation
Use Cases / Scenarios Analysis
Objective: Real-world user detailed descriptions of what is to be supported; narrative may conform to a general template / questions.
Objective: Formal requirements modeling (including “relations” mapping) that forms the basis for bi-directional traceability + testing & tooling
Objective: Standard computable use case / scenario representation. Basis for test tooling & transformation to ReqIF specifications.
Objective: Analyze the “business case” narratives into use cases, realization scenarios & technical use cases (using standardized templates)
Note: See ReqIF background information below.

## Slide 38: Examples for tracing Narrative to Plug-n-Trust
Exemplar considerations from Quiet Hospital Narratives:
  Identify (4) scenario-based requirements that can be verified during CA test
  Include (4) alert priorities: Event + High/Medium/Low Alarms
  Trace through the (4) KIP standards
  Include both Physiological & Technical Alarms
  Include at least (2) modalities + modality-specific alarms (e.g., AIL)
  Include DIS, DAS & C-DAS support, scaling from simple to complex
  Include both Quiet & Silent Hospital requirements
  Include both SDC and FHIR endpoints
Note: SDC alert delegation integrates all four interoperability purposes & SDPi profiles

## Slide 39
What if …
We mashed together the
MDI & SES
Communities & Standards!?!

## Slide 40: A Framework for Trusted Interoperable Product Decoupling
Addressing the SES MDI Ecosystem “Trust Gap” …
SES
MDI Trust Gap!
Diagram from 81001-1:2020 under Open Government License: see https://www.nationalarchives.gov.uk/doc/open-government-licence/version/3/

## Slide 41: SES MDI – Roadmap Proposal
Develop an IHE-HL7 Gemini SES MDI framework concept white paper (underway):
  See confluence.hl7.org “SDPi+FHIR Paper - SES MDI” page
    “Safe, Effective & Secure MDI Using SDC/SDPi + FHIR” (working title)
    Intent is to think through what would be involved in establishing a standards-based framework for Trusted Interoperable Product Decoupling (as identified in the preceding slides)
    See below for potential content and discussion topics
  White paper intended to be jointly balloted and published by IHE & HL7
  Participation is open to all experts – no organizational membership or $$$ required
Add 2020-June TC/215 WG2 & JWG7 TR PWI for the above white paper
  Anticipate Gemini SDPi+FHIR white paper to be submitted to JWG7 via IHE Liaison “fasttrack” process mid-2020
  White paper will include recommendations to JWG7 for standards coordination and development
Invite WG2 & JWG7 members and parent committee experts to participate in the white paper
Consider recommendations to WG2 & JWG7 later in 2020; may include a proposal for an SES MDI international standard and guidance (e.g., 81001-7-x) leveraging content in the white paper

## Slide 42: SES MDI using SDC-SDPi+FHIR - Content
White Paper content may include:
  Subtitle: Strategy for enabling an ecosystem of trusted interoperable decoupled products
  Consider the Plug-n-Trust “pathway” through the SES Temple + MDI Layers
  Identify and discuss the application of standards to be considered …
    Including coordination with and normative references to/from related standards …
    ISO/IEEE 11073 SDC, HL7 FHIR, JWG7 81001-1, 80001-1/-2X, 82304-X, 62304,…)
    In the parent committees (e.g., 60601-2-x & 11073 family)
    Other ISO/IEC committees (e.g., JTC1, ISO TC210)
  SES MDI computable assurance case concept, aligned with related JWG7 work
  Address the use of OMG ReqIF to represent both MDI and SES requirements
  Consider how to integrate ReqIF specifications + Plug-n-Trust SFC CA results into computable assurance cases
  Strategy and roadmap recommendations for addressing standards gaps, implementation and test tooling, and multi-stakeholder collaboration & coordination, etc., necessary to address the …
SES MDI Trust Gap!

## Slide 43
Additional
Information

## Slide 44
OMG Requirements Interchange Format (ReqIF) “Mapping”
Background

## Slide 45: OMG ReqIF: Base Model

## Slide 46: ReqIF: Requirement Hierarchies & Relationships
“mapping” between
Source & Target Requirements