## Slide 1: DIS vs. DAS vs. CDAS – v3
DIS is a system for reporting alarm signals with no technical confirmation (of receipt).
  Cannot rely on it for alarm signaling as a risk control
  Optional support operator alarm management* response locally
    Example - patient remote display, hallway display, one-way pager
CDIS is a system for reporting alarm signals with no technical confirmation and operator confirmation (accept/reject).
  Cannot rely on it for alarm signaling as a risk control
  Optional support operator alarm management* response locally and remotely
    Example – two-way pager (open loop)
DAS is a system for reporting alarm signals with technical confirmation (of receipt).
  Can rely on it for alarm signaling as a risk control
  Optional support operator alarm management* response locally
  A communications failure or failure in any remote component of the DAS must initiate a technical alarm.
    Example - Central Station
CDAS is a system for reporting alarm signals with technical and operator confirmation (accept/reject) (of receipt).
  Can rely on it for alarm signaling as a risk control
  Supports operator confirmation (accept/reject); It may redirect…
  Optional support operator alarm management* response both locally and remotely
  A communications failure or failure in any remote component of the DAS must initiate a technical alarm.
    Example - System that sends alarm to caregiver mobile device with accept / reject. Integrator may redirect
* Alarm management is the ability to inactivate the alarm in some fashion

## Slide 2: Alert System Variants – v3
Description | Type | Technical Delivery Confirmation* | Operator Delivery Confirmation** | Optional Alarm Management | Example
Reports alerts from devices for a Single Patient | DISsp | No | No | Local | Single-Pt. information dashboard
 | CDISsp | No | Yes*** | Remote*** | Single-Pt. information dashboard w/accept/reject
 | DASsp | Yes | No | Local | Single-Pt. Remote View w/audible alarms
 | CDASsp | Yes | Yes | Remote | Single-Pt. Remote View w/accept/reject
 | DISmp | No | No | Local | Multi-Pt. info. dashboard or View Station
Reports alerts for Multiple Patients | CDISmp | No | Yes*** | Remote*** | Multi-Pt. info. View Station w/accept/reject
 | DASmp | Yes | No | Local | Multi-Pt. View Station with audible alarms
 | CDASmp | Yes | Yes | Remote | Multi-Pt. Central Station w/accept/reject
Reports and directs alerts to a responsible CareGiver | DIScg | No | No | Local | Alerts to caregiver pager, Mobile viewer w/accept/reject
 | CDIScg | No | Yes*** | Remote*** | Alerts to caregiver pager, Mobile viewer
 | DAScg | Yes | No | Local | Alerts to caregiver w/audible/haptic alarms
 | CDAScg | Yes | Yes | Remote | Alerts to caregiver w/accept/reject
* In each communication step the receiving device provides a technical response to the sending device that it received and is taking responsibility of the alert
**Operator can, at their choice, use the receiving device (communicator) UI to accept or reject responsibility for the alert
*** Not recommended since there is no confirmation that the Source has received the commands

## Slide 3: Summary
Quiet Room is possible with:
  CDISmp after Accept
    Receiving the operator Accept the source device may decide to enter audio pause
Silent Room is possible with:
  DASmp or CDASmp
    If alerts are delegated to DASmp or CDASmp
Silent ICU is possible with:
  DAScg or CDAScg
    If alerts are delegated to DAScg or CDAScg
    If caregiver device is treated as alert risk control